## Architecture
Block(5) | 목적 | 핵심 구성(부품/모듈) | 주요 노드/배선 포인트 | PoC Bring-up 체크
PC | 데이터 로깅/제어 UI. USB로 MCU 전원+통신. | 18: Fuse holder, 5x20mm PCB clip (pair) (732-11373-ND) 33: DC barrel jack, 5.5mm OD x 2.5mm ID, PCB right-angle, through-hole (SC1863-ND) | USB 케이블(PC↔Arduino). | Serial 통신 확인(예: 'hello' 프린트).
MCU | SPI로 ADC 읽기, 릴레이/전류원 제어, 계산(Rs=V/I). | 1: Arduino Nano 33 IoT (3.3V) (1050-ABX00027-ND) | SPI: SCK↔CLK, MOSI↔DIN, MISO↔DOUT, CS#, DRDY#. 3V3/GND는 스타에. | ADS1220 ID/레지스터 읽기, DRDY 인터럽트(옵션) 확인.
Power | 12V 공급 및 보호(퓨즈/TVS/벌크). 3V3는 Arduino USB에서. | 15: 12V wall adapter, 60W (12V/5A), 2.5mm center-positive (1939-1237-ND) 16: Power Cord (TL868-ND) 17: Fuse, 5x20mm, 500mA slow-blow (0215.500MXEP-ND) 19: Toggle switch SPST, panel mount (2057-SW-T3-1A-A-A3-S1-ND) 20: Power LED, 3mm red through-hole (4491-LED3RED-ND) 34: Inline blade fuse holder (ATC/ATO), splashproof cover, wire leads (F3209-ND) 35: ATC/ATO blade fuse, 5A, 32V (0287005.U-ND) 36: ATC/ATO blade fuse, 3A, 32V (F11061-ND) 37: TVS diode, SMBJ15A, unidirectional, DO-214AA (SMBJ15ALFCT-ND) 38: Bulk electrolytic capacitor, low ESR, 470uF 25V radial (P14418-ND) | 12V 입력→(퓨즈)→분기(모터/4PP). TVS는 퓨즈 뒤 12V↔GND 병렬. 470µF는 12V 근처. | 무부하 12V 확인→퓨즈 후단 전압 확인→릴레이 코일 구동 시 전압 강하 관찰.
Current | 정전류 인가(예: 1mA/100µA), (옵션) 전류 모니터/전류 반전. | 3: Adjustable current source IC (TO-92) (LM334ZNS/NOPB-ND) 4: Rset for 1.0 mA (68.1Ω, 0.1%) (1712-YR1B68R1CCCT-ND) 5: Rset for 100 µA (681Ω, 0.1%) (1712-YR1B681RCCCT-ND) 6: Shunt resistor for current monitor (100Ω, 0.1%) (13-MFP-25BRD52-100RCT-ND) 7: DPDT relay, 12V coil (through-hole) (255-1002-5-ND) 8: NPN transistor (TO-92) (PN2222AFS-ND) 9: Flyback diode for relay coil (1N4148FS-ND) 10: Base resistor (1kΩ, 0.1% optional; 1% OK) (A105942CT-ND) | LM334+Rset로 I 설정. (옵션) 100Ω shunt 직렬 후 전압을 ADC로 읽어 I 계산. DPDT 릴레이로 I 방향 반전(옵션). | 더미 저항(예: 1k~10k) 연결 후 shunt 전압으로 실제 전류 확인.
Voltage | V센스(내측 프로브) 차동 전압을 24-bit ADC로 측정. | 2: ADS1220 24-bit ADC breakout (1188-BB-ADS1220-ND) 11: Diode Array 1 Pair Common Anode 30 V 200mA Surface Mount TO-236-3, SC-59, SOT-23-3 (4878-BAT54ACT-ND) 12: Diode Array 1 Pair Common Cathode 30 V 200mA Surface Mount TO-236-3, SC-59, SOT-23-3 (4878-BAT54CTR-ND) 13: SOT23 TO DIP ADAPTER (1568-00717-ND) 14: Series resistors for ADC inputs (1kΩ, 1%) (1.00KXBK-ND) 21: Electrolytic capacitor 10µF 25V radial (through-hole) (P15799CT-ND) 22: Film capacitor 0.1µF 50V radial (through-hole) (399-5453-1-ND) 39: 10 Ohms ±1% 0.25W, 1/4W Through Hole Resistor Axial Metal Film (13-MFR-25FTF52-10RCT-ND) 40: 1 µF ±10% 16V Ceramic Capacitor X7R Radial (445-FK24X7R1C105KN000-ND) 41: 0.1 µF ±10% 50V Ceramic Capacitor X7R Radial (490-9144-1-ND) | IN0=V+, IN1=V-. 각 입력 앞 1k 직렬. BAT54A/C로 레일 클램프. 3V3_A 레일은 10Ω 뒤쪽 + 0.1µF¦¦1µF(3V3_A↔AGND). 실드는 ADC쪽에서만 접지. | 입력 쇼트(IN0=IN1) 시 ~0V, 작은 배터리/표준저항으로 mV 레벨 변화 확인.
Support | 배선/커넥터/기구/소모품 등 PoC 완성도 향상. | 23: Male header, 1x40, 2.54mm, through-hole (HDR100IMP40M-G-V-TH-ND) 24: Female header, 1x40, 2.54mm, through-hole (HDR100IMP40F-G-V-TH-ND) 25: CONN BIND POST KNURLED BLACK (501-1094-ND) 26: Binding Post Connector Standard Knurled Red Tin (501-1095-ND) 27: Binding Post Connector Standard Knurled Green Tin (501-1506-ND) 28: CONN BIND POST KNURLED BLUE (501-1650-ND) 29: Ferrite clamp (snap-on) for motor cable (1934-1375-ND) 30: shielded wire 22AWG (for V+ / V- lines, optional) (135-94610605000-DS-ND) 31: Desolder braid/wick (for fixing mistakes) (60-5-5-ND) 32: SWITCH SNAP ACTION SPDT 15A 250V (SW1249-ND) 42: 2-pin 0.1" pitch terminal block (optional convenience) (Product ID 2138) 43: Micro-USB cable (Nano 33 IoT to PC) (Product ID 2185) ... +19개 | 바인딩포스트(패널)→내부 배선→보드. 스트레인 릴리프. Belden 9461은 V+/V-용(실드는 ADC쪽만). | 케이블 당김 시 납땜부 스트레스 없는지, 커넥터 핀아웃 라벨링 확인.

## Sheet1
Item # | Category | Block(5) | Subsystem | Item | Qty | Supplier | Supplier_PN | Mfr | Mfr_PN | Specs/Notes | URL | Unit Price | Total | Notes (misc)
1 | MCU/제어 | MCU | MCU/제어 | Arduino Nano 33 IoT (3.3V) | 3 | Digi-Key | 1050-ABX00027-ND | Arduino | ABX00027 | 3.3V logic (no level shifting needed with ADS1220). USB-powered. | https://www.digikey.com/en/products/detail/arduino/ABX00027/10239968 | 23.9
2 | 전압측정/ADC | Voltage | ADC/전압측정 | ADS1220 24-bit ADC breakout | 3 | Digi-Key | 1188-BB-ADS1220-ND | Olimex LTD | BB-ADS1220 | 24-bit ΔΣ ADC w/ PGA; SPI. Measures V(2-3) and also shunt current if desired. | https://www.digikey.com/en/products/detail/olimex-ltd/BB-ADS1220/21662042 | 18.64
3 | 정전류/전류모니터 | Current | 정전류 인가 | Adjustable current source IC (TO-92) | 10 | Digi-Key | LM334ZNS/NOPB-ND | Texas Instruments | LM334Z/NOPB | Sets current via Rset: I≈67.7mV/Rset (@25°C). Use with 12V supply. | https://www.digikey.com/en/products/detail/texas-instruments/LM334Z-NOPB/6215 | 0.637
4 | 정전류/전류모니터 | Current | 정전류 인가 | Rset for 1.0 mA (68.1Ω, 0.1%) | 25 | Digi-Key | 1712-YR1B68R1CCCT-ND | TE Connectivity Passive Product | YR1B68R1CC | I≈0.0677/68.1≈0.994 mA at 25°C. Axial (easy solder). | https://www.digikey.com/en/products/detail/te-connectivity-passive-product/YR1B68R1CC/2390921 | 0.4628
5 | 정전류/전류모니터 | Current | 정전류 인가 | Rset for 100 µA (681Ω, 0.1%) | 25 | Digi-Key | 1712-YR1B681RCCCT-ND | TE Connectivity Passive Product | YR1B681RCC | I≈0.0677/681≈99.4 µA at 25°C. Axial (easy solder). | https://www.digikey.com/en/products/detail/te-connectivity-passive-product/YR1B681RCC/2390920 | 0.4628
6 | 정전류/전류모니터 | Current | 정전류 인가 | Shunt resistor for current monitor (100Ω, 0.1%) | 25 | Digi-Key | 13-MFP-25BRD52-100RCT-ND | YAGEO | MFP-25BRD52-100R | Place in series with current; read Vshunt with ADS1220 to compute I. | https://www.digikey.com/en/products/detail/yageo/MFP-25BRD52-100R/2058796 | 0.3272
7 | 정전류/전류모니터 | Current | 정전류 인가 | DPDT relay, 12V coil (through-hole) | 3 | Digi-Key | 255-1002-5-ND | Panasonic Electric Works | TQ2-12V | Swap I+ / I- to do +I/-I measurements; improves stability. | https://www.digikey.com/en/products/detail/panasonic-electric-works/TQ2-12V/251774 | 2.16
8 | 정전류/전류반전(옵션) | Current | 전류 반전/릴레이 | NPN transistor (TO-92) | 10 | Digi-Key | PN2222AFS-ND | Onsemi | PN2222ABU | replace PN2222A | https://www.digikey.com/en/products/detail/onsemi/pn2222abu/6534 | 0.184
9 | 정전류/전류반전(옵션) | Current | 전류 반전/릴레이 | Flyback diode for relay coil | 20 | Digi-Key | 1N4148FS-ND | Onsemi | 1N4148FS-ND | replace 1N4148-T | https://www.digikey.com/en/products/detail/onsemi/1N4148/458603 | 0.41
10 | 정전류/전류반전(옵션) | Current | 전류 반전/릴레이 | Base resistor (1kΩ, 0.1% optional; 1% OK) | 25 | Digi-Key | A105942CT-ND | TE Connectivity Passive Product | YR1B1K0CC | Use 1kΩ from Arduino pin to PN2222A base. | https://www.digikey.com/en/products/detail/te-connectivity-passive-product/YR1B1K0CC/2390772 | 0.4568
11 | 전압측정/ADC | Voltage | ADC/전압측정 | Diode Array 1 Pair Common Anode 30 V 200mA Surface Mount TO-236-3, SC-59, SOT-23-3 | 10 | Digi-Key | 4878-BAT54ACT-ND | Diotec Semiconductor | BAT54A | BAT54A(공통 애노드): 입력 하단 클램프. pin3(공통)→AGND, pin1→IN0 노드, pin2→IN1 노드. (각 입력 앞에 1k 직렬저항 권장). SOT-23이라 00717 어댑터 사용. | https://www.digikey.com/en/products/detail/diotec-semiconductor/BAT54A/13163440 | 0.088
12 | 전압측정/ADC | Voltage | ADC/전압측정 | Diode Array 1 Pair Common Cathode 30 V 200mA Surface Mount TO-236-3, SC-59, SOT-23-3 | 10 | Digi-Key | 4878-BAT54CTR-ND | Diotec Semiconductor | BAT54C | BAT54C(공통 캐소드): 입력 상단 클램프. pin3(공통)→3V3_A, pin1→IN0 노드, pin2→IN1 노드. (각 입력 앞에 1k 직렬저항 권장). SOT-23이라 00717 어댑터 사용. | https://www.digikey.com/en/products/detail/diotec-semiconductor/BAT54C/13163831 | 0.092
13 | 전압측정/ADC | Voltage | ADC/전압측정 | SOT23 TO DIP ADAPTER | 10 | Digi-Key | 1568-00717-ND | SparkFun Electronics | 00717 | SOT-23→DIP 변환 보드. BAT54A용 1개 + BAT54C용 1개 필요(=최소 2개). SOT-23 핀3(공통) 매핑은 멀티미터 연속(삐-)로 확인 권장. | https://www.digikey.com/en/products/detail/sparkfun-electronics/00717/5318740 | 1.25
14 | 전압측정/ADC | Voltage | ADC/전압측정 | Series resistors for ADC inputs (1kΩ, 1%) | 50 | Digi-Key | 1.00KXBK-ND | YAGEO | MFR-25FBF52-1K | 1kΩ axial 1% (through-hole). Put in series with V+ and V- into ADC to limit current during faults. | https://www.digikey.com/en/products/detail/yageo/MFR-25FBF52-1K/13011 | 0.0286
15 | 전원/보호 | Power | 전원/보호 | 12V wall adapter, 60W (12V/5A), 2.5mm center-positive | 1 | Digi-Key | 1939-1237-ND | GlobTek, Inc. | TR9CE5000LCP-N(R6B) | Pleading simple external power. USB powers Arduino; tie grounds between USB GND and 12V GND. | https://www.digikey.com/en/products/detail/globtek-inc/TR9CE5000LCP-N-R6B/8598040 | 19.24
16 | 전원/보호 | Power | 전원/보호 | Power Cord | 1 | Digi-Key | TL868-ND | Eaton Tripp Lite | P006-002-13A | IEC C13 power cord (for IEC C14 inlet) | https://www.digikey.com/en/products/detail/eaton-tripp-lite/P006-002-13A/4896253?s=N4IgTCBcDaIAoAYEDYC0SyoIwGYCCIAugL5A | 8.23
17 | 전원/보호 | Power | 전원/보호 | Fuse, 5x20mm, 500mA slow-blow | 10 | Digi-Key | 0215.500MXEP-ND | Littelfuse Inc. | 0215.500MXEP | Optional but good practice on the 12V input. | https://www.digikey.com/en/products/detail/littelfuse-inc/0215-500MXEP/777421 | 2.275
18 | PC/호스트 | PC | PC/연결 | Fuse holder, 5x20mm PCB clip (pair) | 5 | Digi-Key | 732-11373-ND | Würth Elektronik | 696105003002 | Two clips per fuse (buy 2 fuses worth = 4 clips). Alternatively use a panel fuse holder. | https://www.digikey.com/en/products/detail/w%C3%BCrth-elektronik/696105003002/7244557 | 0.61
19 | 전원/보호 | Power | 전원/보호 | Toggle switch SPST, panel mount | 3 | Digi-Key | 2057-SW-T3-1A-A-A3-S1-ND | Adam Tech | SW-T3-1A-A-A3-S1 | Optional on/off for 12V input. | https://www.digikey.com/en/products/detail/adam-tech/SW-T3-1A-A-A3-S1/15284460 | 0.98
20 | 전원/보호 | Power | 전원/보호 | Power LED, 3mm red through-hole | 50 | Digi-Key | 4491-LED3RED-ND | Lumimax Optoelectronic Technology | LED3RED | Optional power indicator. | https://www.digikey.com/en/products/detail/lumimax-optoelectronic-technology/LED3RED/24389652 | 0.073
21 | 전압측정/전원디커플링 | Voltage | ADC 전원/노이즈 | Electrolytic capacitor 10µF 25V radial (through-hole) | 10 | Digi-Key | P15799CT-ND | Panasonic Electronic Components | EEA-GA1E100H | Across 12V rail near relay/current source. | https://www.digikey.com/en/products/detail/panasonic-electronic-components/EEA-GA1E100H/2513301 | 0.173
22 | 전압측정/전원디커플링 | Voltage | ADC 전원/노이즈 | Film capacitor 0.1µF 50V radial (through-hole) | 10 | Digi-Key | 399-5453-1-ND | KEMET | R82EC3100DQ70J | replacing ECQ-B1H104JF3 | https://www.digikey.com/en/products/detail/kemet/R82EC3100DQ70J/1930807?s=N4IgTCBcDaIMwE4EFoCsAWVdkEZkDkAREAXQF8g | 0.183
23 | 지원/배선·커넥터 | Support | 배선/커넥터 | Male header, 1x40, 2.54mm, through-hole | 10 | Digi-Key | HDR100IMP40M-G-V-TH-ND | Chip Quik Inc. | HDR100IMP40M-G-V-TH | Solder to Olimex board / perfboard as needed. | https://www.digikey.com/en/products/detail/chip-quik-inc/HDR100IMP40M-G-V-TH/5978197 | 0.69
24 | 지원/배선·커넥터 | Support | 배선/커넥터 | Female header, 1x40, 2.54mm, through-hole | 10 | Digi-Key | HDR100IMP40F-G-V-TH-ND | Chip Quik Inc. | HDR100IMP40F-G-V-TH | Make sockets so boards are pluggable; cut to length. | https://www.digikey.com/en/products/detail/chip-quik-inc/HDR100IMP40F-G-V-TH/5978200 | 1.09
25 | 지원/배선·커넥터 | Support | 배선/커넥터 | CONN BIND POST KNURLED BLACK | 2 | Digi-Key | 501-1094-ND | Pomona Electronics | 3760-0 | For I-/V- (or your labeling). | https://www.digikey.com/en/products/detail/pomona-electronics/3760-0/603401 | 6.29
26 | 지원/배선·커넥터 | Support | 배선/커넥터 | Binding Post Connector Standard Knurled Red Tin | 2 | Digi-Key | 501-1095-ND | Pomona Electronics | 3760-2 | For I+/V+ (or your labeling). | https://www.digikey.com/en/products/detail/pomona-electronics/3760-2/603402 | 6.29
27 | 지원/배선·커넥터 | Support | 배선/커넥터 | Binding Post Connector Standard Knurled Green Tin | 2 | Digi-Key | 501-1506-ND | Pomona Electronics | 3760-5 | For I-/V- (or your labeling). | https://www.digikey.com/en/products/detail/pomona-electronics/3760-5/736554 | 6.29
28 | 지원/배선·커넥터 | Support | 배선/커넥터 | CONN BIND POST KNURLED BLUE | 2 | Digi-Key | 501-1650-ND | Pomona Electronics | 3760-6 | For I+/V+ (or your labeling). | https://www.digikey.com/en/products/detail/pomona-electronics/3760-6/736555 | 6.29
29 | 지원/기타 | Support | 기타 | Ferrite clamp (snap-on) for motor cable | 10 | Digi-Key | 1934-1375-ND | Fair-Rite | 431167281 | reduce stepper cable EMI that can leak into 4PP ADC. Place near driver and/or near motor. | https://www.digikey.com/en/products/filter/ferrite-cores/933 | 1.792
30 | 지원/배선·커넥터 | Support | 배선/커넥터 | shielded wire 22AWG (for V+ / V- lines, optional) | 25 | Digi-Key | 135-94610605000-DS-ND | Belden Inc. | 9461 0605000 | Optional: use shield cable for voltage sense lines. If you skip, twist V+ and V- tightly. | https://www.digikey.com/en/products/detail/belden-inc/9461-0605000/8132459?s=N4IgTCBcDaIIwGYCsBaAnAFgGxwAxdyV2JQBEBlFAOVJAF0BfIA | 2.9336
31 | 지원/공구·소모품 | Support | 공구/소모품 | Desolder braid/wick (for fixing mistakes) | 1 | Digi-Key | 60-5-5-ND | Chemtronics | 60-5-5 | 5 ft wick; pick any size you like. | https://www.digikey.com/en/products/detail/chemtronics/60-5-5/306973 | 8.31
32 | 지원/기구(스테이지 제외) | Support | 기구/스테이지(제외) | SWITCH SNAP ACTION SPDT 15A 250V | 10 | Digi-Key | SW1249-ND | Omron | V-153-1A5 | switch for zeroing stepper motors | https://www.digikey.com/en/products/detail/omron-electronics-inc-emc-div/V-153-1A5/1829012 | 3.371
33 | PC/호스트 | PC | PC/연결 | DC barrel jack, 5.5mm OD x 2.5mm ID, PCB right-angle, through-hole | 2 | Digi-Key | SC1863-ND | Switchcraft Inc. | RAPC712BK | 12V power barrel jack | https://www.digikey.com/en/products/detail/switchcraft-inc/RAPC712BK/1288493 | 3.09
34 | 전원/보호 | Power | 전원/보호 | Inline blade fuse holder (ATC/ATO), splashproof cover, wire leads | 5 | Digi-Key | F3209-ND | Littelfuse/Commercial Vehicle Products | FHAC0001ZXJ | 12V 메인 라인에 직렬로 넣어서 전체 보호(권장). 2개는 '메인/모터' 분기 각각 보호할 때 사용. | https://www.digikey.com/en/products/detail/littelfuse-commercial-vehicle-products/FHAC0001ZXJ/2004062 | 5.33
35 | 전원/보호 | Power | 전원/보호 | ATC/ATO blade fuse, 5A, 32V | 10 | Digi-Key | 0287005.U-ND | Littelfuse Inc. | 0287005.U | 12V 5A 어댑터 기준 메인 퓨즈로 무난. (실제 동작 전류가 낮으면 3A로 낮추는 게 더 안전) | https://www.digikey.com/en/products/detail/littelfuse-inc/0287005.U/3996899 | 0.332
36 | 전원/보호 | Power | 전원/보호 | ATC/ATO blade fuse, 3A, 32V | 10 | Digi-Key | F11061-ND | Littelfuse Inc. | 0287003.U | 12V 3A 어댑터 메인 퓨즈로 무난. (실제 동작 전류가 5A보다 낮을때) | https://www.digikey.com/en/products/detail/littelfuse-inc/0287003-U/3996895?s=N4IgTCBcDaIAxgBwHY5wMwDoCqIC6AvkA | 0.332
37 | 전원/보호 | Power | 전원/보호 | TVS diode, SMBJ15A, unidirectional, DO-214AA | 10 | Digi-Key | SMBJ15ALFCT-ND | Littelfuse Inc. | SMBJ15A | 12V 레일에 병렬(12V↔GND)로 넣어 스파이크/ESD 흡수. 퓨즈 '뒤'에 배치 권장. | https://www.digikey.com/en/products/detail/littelfuse-inc/SMBJ15A/285980 | 0.189
38 | 전원/보호 | Power | 전원/보호 | Bulk electrolytic capacitor, low ESR, 470uF 25V radial | 10 | Digi-Key | P14418-ND | Panasonic Electronic Components | EEU-FR1E471L | VMOT(모터 드라이버) 근처 12V↔GND에 추가(권장). 리플/스텝 부하에 대한 버퍼 역할. (Panasonic FR, 105°C) | https://www.digikey.com/en/products/detail/panasonic-electronic-components/EEU-FR1E471L/2433552 | 0.592
39 | 전압측정/전원디커플링 | Voltage | ADC 전원/노이즈 | 10 Ohms ±1% 0.25W, 1/4W Through Hole Resistor Axial Metal Film | 25 | Digi-Key | 13-MFR-25FTF52-10RCT-ND | YAGEO | MFR-25FTF52-10R | 옵션(권장): Arduino 3V3_DIG→10Ω→3V3_A(ADC 전용 전원). 10Ω '뒤쪽(3V3_A)'에서 0.1µF+1µF를 AGND로 디커플링. ADS1220 VDD/AVDD 및 BAT54C 공통(#3)도 3V3_A에 연결. | https://www.digikey.com/en/products/detail/yageo/MFR-25FTF52-10R/9140488?utm_source=chatgpt.com | 0.0312
40 | 전압측정/전원디커플링 | Voltage | ADC 전원/노이즈 | 1 µF ±10% 16V Ceramic Capacitor X7R Radial | 50 | Digi-Key | 445-FK24X7R1C105KN000-ND | TDK Corporation | FK24X7R1C105KN000 | 필수(권장): 3V3_A↔AGND에 0.1µF와 병렬. ADS1220 VDD/AVDD 핀 근처에 배치(리드/배선 최소). | https://www.digikey.com/en/products/detail/tdk-corporation/fk24x7r1c105kn000/1008884 | 0.1738
41 | 전압측정/전원디커플링 | Voltage | ADC 전원/노이즈 | 0.1 µF ±10% 50V Ceramic Capacitor X7R Radial | 50 | Digi-Key | 490-9144-1-ND | Murata Electronics | RDER71H104K0M1H03A | 필수(권장): 3V3_A↔AGND 디커플링(고주파). 1µF와 병렬로 ADS1220 전원 핀 근처에 배치. | https://www.digikey.com/en/products/detail/murata-electronics/RDER71H104K0M1H03A/4772306 | 0.20840
42 | 지원/배선·커넥터 | Support | 배선/커넥터 | 2-pin 0.1" pitch terminal block (optional convenience) | 30 | Adafruit | Product ID 2138 | Adafruit | 2138 | Screw terminals make wiring easier (no crimp tools). Use for power, relay, etc. | https://www.adafruit.com/product/2138 | 0.86
43 | 지원/배선·커넥터 | Support | 배선/커넥터 | Micro-USB cable (Nano 33 IoT to PC) | 5 | Adafruit | Product ID 2185 | Adafruit | 2185 | USB-A to Micro-B cable (2m). Nano 33 IoT uses Micro-USB. | https://www.adafruit.com/product/2185 | 4.95
44 | 지원/기구(스테이지 제외) | Support | 기구/스테이지(제외) | TMC2209 Stepper Motor Driver | 5 | Adafruit | Product ID 6121 | Adafruit | 6121 | TMC2209 stepper driver breakout | https://www.adafruit.com/product/6121 | 8.95
45 | 지원/배선·커넥터 | Support | 배선/커넥터 | Hook-up wire spool set, 22AWG stranded-core, 10 colors, 10 x 25ft (spools) | 1 | Adafruit | Product ID 3175 | Adafruit | 3175 | Recommended main signal/harness wire for 4PP + TMC2206 STEP/DIR/EN/UART/limit switches. Stranded handles movement well. | https://www.adafruit.com/product/3175 | 29.95
46 | 지원/배선·커넥터 | Support | 배선/커넥터 | Hook-up wire spool set, 22AWG solid-core, 10 colors, 10 x 25ft (spools) | 1 | Adafruit | Product ID 3174 | Adafruit | 3174 | Optional: if you prefer very tidy board wiring and lots of color options. Solid-core holds bends well. | https://www.adafruit.com/product/3174 | 29.95
47 | 지원/공구·소모품 | Support | 공구/소모품 | PCB vise / working station | 1 | Adafruit | Product ID 5253 | Adafruit | 5253 | Holds boards while soldering. Any helping-hands works. | https://www.adafruit.com/product/5253 | 29.95
48 | 지원/기구(스테이지 제외) | Support | 기구/스테이지(제외) | lead screw clamp | 1 | Bilda | 3504-0804-2109 | Bilda | 3504-0804-2109 | 3504 Series Lead Screw Clamping Collar | https://www.gobilda.com/3504-series-lead-screw-clamping-collar-8mm-lead-4-start-21mm-od-9mm-length/ | 6.99
49 | 지원/기구(스테이지 제외) | Support | 기구/스테이지(제외) | thrust ball bearing | 2 | Bilda | 1613-0516-0008 | Bilda | 1613-0516-0008 | 8mm ID Thrust Ball Bearing (15.5mm OD, 5mm Thickness) | https://www.gobilda.com/8mm-id-thrust-ball-bearing-15-5mm-od-5mm-thickness/ | 3.99
50 | 지원/기구(스테이지 제외) | Support | 기구/스테이지(제외) | slotted rail 144mm | 1 | Bilda | 1118-0024-0144 | Bilda | 1118-0024-0144 | 1118 Series Open goRAIL® (144mm Length) | https://www.gobilda.com/1118-series-open-gorail-144mm-length/ | 6.59
51 | 지원/기구(스테이지 제외) | Support | 기구/스테이지(제외) | slotted rail 192mm | 1 | Bilda | 1118-0024-0192 | Bilda | 1118-0024-0192 | 1118 Series Open goRAIL® (192mm Length) | https://www.gobilda.com/1118-series-open-gorail-192mm-length/ | 7.79
52 | 지원/기구(스테이지 제외) | Support | 기구/스테이지(제외) | slotted rail 96mm | 1 | Bilda | 1118-0024-0096 | Bilda | 1118-0024-0096 | 1118 Series Open goRAIL® (96mm Length) |  | 5.39
53 | 지원/기구(스테이지 제외) | Support | 기구/스테이지(제외) | flanged ball bearing | 1 | Bilda | 1611-0514-0008 | Bilda | 1611-0514-0008 | 1611 Series Flanged Ball Bearing (8mm ID x 14mm OD, 5mm Thickness) - 2 Pack | https://www.gobilda.com/1611-series-flanged-ball-bearing-8mm-id-x-14mm-od-5mm-thickness-2-pack/ | 3.99
54 | 지원/기구(스테이지 제외) | Support | 기구/스테이지(제외) | lead screw 150mm | 1 | Bilda | 3501-0804-0150 | Bilda | 3501-0804-0150 | 3501 Series Lead Screw (8mm Lead, 4 Start, 150mm Length) | https://www.gobilda.com/3501-series-lead-screw-8mm-lead-4-start-150mm-length/ | 6.99
55 | 지원/기구(스테이지 제외) | Support | 기구/스테이지(제외) | lead screw 250mm | 1 | Bilda | 3501-0804-0250 | Bilda | 3501-0804-0250 | 3501 Series Lead Screw (8mm Lead, 4 Start, 250mm Length) | https://www.gobilda.com/3501-series-lead-screw-8mm-lead-4-start-250mm-length/ | 8.99
56 | 지원/기구(스테이지 제외) | Support | 기구/스테이지(제외) | lead screw 200mm | 1 | Bilda | 3501-0804-0200 | Bilda | 3501-0804-0200 | 3501 Series Lead Screw (8mm Lead, 4 Start, 200mm Length) | https://www.gobilda.com/3501-series-lead-screw-8mm-lead-4-start-200mm-length/ | 7.99
57 | 지원/기구(스테이지 제외) | Support | 기구/스테이지(제외) | lead screw nut | 1 | Bilda | 3503-0804-0038 | Bilda | 3503-0804-0038 | 3503 Series Lead Screw Nut for Open goRAIL® (8mm Lead, 4 Start, 38mm Length) | https://www.gobilda.com/3503-series-lead-screw-nut-for-open-gorail-8mm-lead-4-start-38mm-length/ | 12.99
58 | 지원/기구(스테이지 제외) | Support | 기구/스테이지(제외) | shaft coupler 8x5mm | 1 | Bilda | 4002-0005-0008 | Bilda | 4002-0005-0008 | 4002 Series Flexible Clamping Shaft Coupler (5mm Round Bore to 8mm Round Bore) | https://www.gobilda.com/4002-series-flexible-clamping-shaft-coupler-5mm-round-bore-to-8mm-round-bore/ | 6.99
59 | 지원/기구(스테이지 제외) | Support | 기구/스테이지(제외) | shaft coupler 8x6mm | 1 | Bilda | 4002-0006-0008 | Bilda | 4002-0006-0008 | 4002 Series Flexible Clamping Shaft Coupler (6mm Round Bore to 8mm Round Bore) | https://www.gobilda.com/4002-series-flexible-clamping-shaft-coupler-6mm-round-bore-to-8mm-round-bore/ | 6.99
60 | 지원/기구(스테이지 제외) | Support | 기구/스테이지(제외) | Stepper Motor | 2 | Stepperonline | 17HS08-1004S | Stepperonline | 17HS08-1004S | Nema 17 Bipolar Stepper Motor 1.8deg 16Ncm(22.66oz.in) 1A Extruder Motor 42x42x21.5mm 4 Wires |  | 7.99
61 | 지원/기구(스테이지 제외) | Support | 기구/스테이지(제외) | Brass Tapered Heat-Set Inserts for Plastic | 1 | McMaster | 94180A331 | McMaster-Carr | 94180A331 | M3 x 0.5 mm Thread Size, 3.8 mm Installed Length | https://www.mcmaster.com/94180A331/ | 20.44
62 | 지원/기구(스테이지 제외) | Support | 기구/스테이지(제외) | Brass Tapered Heat-Set Inserts for Plastic | 1 | McMaster | 94180A351 | McMaster-Carr | 94180A351 | M4 x 0.7 mm Thread Size, 4.7 mm Installed Length | https://www.mcmaster.com/94180A351/ | 21.82
subtotal